## Sheet1
Comparing Operational and Spare trigger values
2017-03-22
| Property |  | Value | Unit
Cycle/ B-train parameters | Resonant frequency |  | 1.8308 | Mhz
| Pre-injection ramp duration |  | 0.3 | s
| Bdot |  | 1513 | G/s
| B-level at marker |  | 0.043 | T
| Current level at marker |  | 33.44 | A
| Transfer function |  | 12.8589 | G/A
Trigger Time | Standard dev of trigger time (operational) |  | 3.7117e-05 | s
| Standard dev of trigger time (spare) |  | 3.636e-05 | s
| Average difference between operational and spare trigger |  | 9.4337e-05 | s
| Standard dev |  | 5.1959e-05
| dt between operational and spare |  | 9.4337e-05 | s
| Gauss difference between oper and spare |  | 0.14273 | G
| Equivalent frequency |  | 607.6996 | Hz
Standard deviation calculated using Current | OPER |  | 0.00286 | A
|  |  | 0.036776 | G
| SPARE |  | 0.0030122 | A
|  |  | 0.038733 | G
Standard deviation calculated using time | OPER |  | 0.037117 | ms
|  |  | 0.056157 | G
| SPARE |  | 0.03636 | ms
|  |  | 0.055013 | G
Settings | Sample size |  | 80 | cycles
| Window delay length |  | 325 | ms
| Trigger threshold level |  | -1.5 | V
Difference between trigger signal and peak of NMR signal | Standard dev of trigger time (operational) |  | 3.7117e-05 | s
| Standard dev of peak time |  | 4.9367e-05 | s
| Average difference between trigger and peak time |  | 0.00011226 | s
| Standard dev |  | 6.1764e-05 | s
| Standard dev of trigger time (spare) |  | 3.636e-05 | s
| Standard dev of peak time |  | 4.0641e-05 | s
| Average difference between trigger and peak time |  | 0.00011081 | s
| Standard dev |  | 5.4532e-05 | s
Difference between trigger signal and peak of NMR signal | Standard dev of trigger time (operational) |  | 3.7117e-05 | s
| Standard dev of peak time |  | 4.3384e-05 | s
| Average difference between trigger and peak time |  | 0.0001083 | s | 0.16386
| Standard dev |  | 5.7095e-05 | s | 0.086385
PARABOLA FIT | Standard dev of trigger time (spare) |  | 3.636e-05 | s
| Standard dev of peak time |  | 4.0981e-05 | s
| Average difference between trigger and peak time |  | 0.00011982 | s | 0.1813
| Standard dev |  | 5.4786e-05 | s | 0.082891