## Sheet1
Comparing Operational and Spare trigger values
2017-03-22
| Property |  | Value | Unit
Cycle/ B-train parameters | Resonant frequency |  | 14.4759 | Mhz
| Ramp duration |  | 1.5 | s
| Bdot |  | 1950 | G/s
| B-level at marker |  | 0.34 | T
| Current level at marker |  | 258 | A
| Transfer function |  | 13.1783 | G/A
Trigger Time | Standard dev of trigger time (operational) |  | 1.3837e-05 | s
| Standard dev of trigger time (spare) |  | 1.3691e-05 | s
| Average difference between operational and spare trigger |  | 0.000424 | s
| Standard dev |  | 1.9466e-05
| dt between operational and spare |  | 0.000424 | s
| Gauss difference between oper and spare |  | 0.8268 | G
| Equivalent frequency |  | 3520.1903 | Hz
Standard deviation calculated using time | OPER |  | 0.013837 | ms
|  |  | 0.026982 | G
| SPARE |  | 0.013691 | ms
|  |  | 0.026697 | G
Settings | Sample size |  | 100 | cycles
| Window delay length |  | 2769 | ms
| Trigger threshold level |  | -1.5 | V
Difference between trigger signal and peak of NMR signal | Standard dev of trigger time (operational) |  | 1.3837e-05 | s
| Standard dev of peak time |  | 1.3825e-05 | s
| Average difference between trigger and peak time |  | 1.296e-05 | s | 0.025273
| Standard dev |  | 1.956e-05 | s | 0.038142
| Standard dev of trigger time (spare) |  | 1.3691e-05 | s
| Standard dev of peak time |  | 1.35e-05 | s
| Average difference between trigger and peak time |  | 1.3149e-05 | s | 0.025641
| Standard dev |  | 1.9227e-05 | s | 0.037493
Difference between trigger signal and peak of NMR signal | Standard dev of trigger time (operational) |  | 1.3837e-05 | s
| Standard dev of peak time |  | 1.4015e-05 | s
| Average difference between trigger and peak time |  | 1.3109e-05 | s | 0.025563
| Standard dev |  | 1.9695e-05 | s | 0.038404
PARABOLA FIT | Standard dev of trigger time (spare) |  | 1.3691e-05 | s
| Standard dev of peak time |  | 1.325e-05 | s
| Average difference between trigger and peak time |  | 1.2653e-05 | s | 0.024673
| Standard dev |  | 1.9053e-05 | s | 0.037153