## Sheet1
Comparing Operational and Spare trigger values
2017-03-22
| Property |  | Value | Unit
Cycle/ B-train parameters | Resonant frequency |  | 1.8946 | Mhz
| Pre-injection ramp duration |  | 0.3 | s
| Bdot |  | 1513 | G/s
| B-level at marker |  | 0.0445 | T
| Current level at marker |  | 33.58 | A
| Transfer function |  | 13.2519 | G/A
Trigger Time | Standard dev of trigger time (operational) |  | 4.666e-05 | s
| Standard dev of trigger time (spare) |  | 4.6993e-05 | s
| Average difference between operational and spare trigger |  | 0.0001635 | s
| Standard dev |  | 6.6223e-05
| dt between operational and spare |  | 0.0001635 | s
| Gauss difference between oper and spare |  | 0.24738 | G
| Equivalent frequency |  | 1053.2279 | Hz
Standard deviation calculated using Current | OPER |  | 0.0030089 | A
|  |  | 0.039874 | G
| SPARE |  | 0.0025224 | A
|  |  | 0.033427 | G
Standard deviation calculated using time | OPER |  | 0.04666 | ms
|  |  | 0.070597 | G
| SPARE |  | 0.046993 | ms
|  |  | 0.071101 | G
Settings | Sample size |  | 78 | cycles
| Window delay length |  | 325 | ms
| Trigger threshold level |  | -1.5 | V
Difference between trigger signal and peak of NMR signal | Standard dev of trigger time (operational) |  | 4.666e-05 | s
| Standard dev of peak time |  | 6.1046e-05 | s
| Average difference between trigger and peak time |  | 0.0001211 | s
| Standard dev |  | 7.6836e-05 | s
| Standard dev of trigger time (spare) |  | 4.6993e-05 | s
| Standard dev of peak time |  | 4.8739e-05 | s
| Average difference between trigger and peak time |  | 0.00012671 | s
| Standard dev |  | 6.7704e-05 | s
Difference between trigger signal and peak of NMR signal | Standard dev of trigger time (operational) |  | 4.666e-05 | s
| Standard dev of peak time |  | 5.2807e-05 | s
| Average difference between trigger and peak time |  | 0.00011744 | s | 0.17768
| Standard dev |  | 7.0468e-05 | s | 0.10662
PARABOLA FIT | Standard dev of trigger time (spare) |  | 4.6993e-05 | s
| Standard dev of peak time |  | 4.7267e-05 | s
| Average difference between trigger and peak time |  | 0.00012656 | s | 0.19149
| Standard dev |  | 6.6652e-05 | s | 0.10085