## Sheet1
Comparing Operational and Spare trigger values
2017-03-22
| Property |  | Value | Unit
Cycle/ B-train parameters | Resonant frequency |  | 1.9125 | Mhz
| Pre-injection ramp duration |  | 0.3 | s
| Bdot |  | 1513 | G/s
| B-level at marker |  | 0.045 | T
| Current level at marker |  | 33.9 | A
| Transfer function |  | 13.2743 | G/A
Trigger Time | Standard dev of trigger time (operational) |  | 5.4059e-05 | s
| Standard dev of trigger time (spare) |  | 5.9154e-05 | s
| Average difference between operational and spare trigger |  | 0.00022704 | s
| Standard dev |  | 8.0135e-05
| dt between operational and spare |  | 0.00022704 | s
| Gauss difference between oper and spare |  | 0.34351 | G
| Equivalent frequency |  | 1462.5374 | Hz
Standard deviation calculated using Current | OPER |  | 0.0025277 | A
|  |  | 0.033554 | G
| SPARE |  | 0.0028358 | A
|  |  | 0.037643 | G
Standard deviation calculated using time | OPER |  | 0.054059 | ms
|  |  | 0.081791 | G
| SPARE |  | 0.059154 | ms
|  |  | 0.0895 | G
Settings | Sample size |  | 100 | cycles
| Window delay length |  | 1000 | ms
| Trigger threshold level |  | -1.5 | V
Difference between trigger signal and peak of NMR signal | Standard dev of trigger time (operational) |  | 5.4059e-05 | s
| Standard dev of peak time |  | 5.9927e-05 | s
| Average difference between trigger and peak time |  | 0.00012417 | s
| Standard dev |  | 8.0707e-05 | s
| Standard dev of trigger time (spare) |  | 5.9154e-05 | s
| Standard dev of peak time |  | 6.1554e-05 | s
| Average difference between trigger and peak time |  | 0.0001333 | s
| Standard dev |  | 8.537e-05 | s
Difference between trigger signal and peak of NMR signal | Standard dev of trigger time (operational) |  | 5.4059e-05 | s
| Standard dev of peak time |  | 5.5747e-05 | s
| Average difference between trigger and peak time |  | 0.00012176 | s | 0.18422
| Standard dev |  | 7.7654e-05 | s | 0.11749
PARABOLA FIT | Standard dev of trigger time (spare) |  | 5.9154e-05 | s
| Standard dev of peak time |  | 6.2672e-05 | s
| Average difference between trigger and peak time |  | 0.00014861 | s | 0.22485
| Standard dev |  | 8.618e-05 | s | 0.13039