## Sheet1
Comparing Operational and Spare trigger values
2017-03-22
| Property |  | Value | Unit
Cycle/ B-train parameters | Resonant frequency |  | 1.9125 | Mhz
| Pre-injection ramp duration |  | 0.6 | s
| Bdot |  | 757 | G/s
| B-level at marker |  | 0.045 | T
| Current level at marker |  | 33.89 | A
| Transfer function |  | 13.2783 | G/A
Trigger Time | Standard dev of trigger time (operational) |  | 7.5736e-05 | s
| Standard dev of trigger time (spare) |  | 7.4877e-05 | s
| Average difference between operational and spare trigger |  | 0.00029275 | s
| Standard dev |  | 0.0001065
| dt between operational and spare |  | 0.00029275 | s
| Gauss difference between oper and spare |  | 0.22161 | G
| Equivalent frequency |  | 943.536 | Hz
Standard deviation calculated using Current | OPER |  | 0.0028332 | A
|  |  | 0.03762 | G
| SPARE |  | 0.0027212 | A
|  |  | 0.036132 | G
Standard deviation calculated using time | OPER |  | 0.075736 | ms
|  |  | 0.057332 | G
| SPARE |  | 0.074877 | ms
|  |  | 0.056682 | G
Settings | Sample size |  | 100 | cycles
| Window delay length |  | 625 | ms
| Trigger threshold level |  | -1.5 | V
Difference between trigger signal and peak of NMR signal | Standard dev of trigger time (operational) |  | 7.5736e-05 | s
| Standard dev of peak time |  | 7.7885e-05 | s
| Average difference between trigger and peak time |  | 0.00014164 | s
| Standard dev |  | 0.00010864 | s
| Standard dev of trigger time (spare) |  | 7.4877e-05 | s
| Standard dev of peak time |  | 8.1356e-05 | s
| Average difference between trigger and peak time |  | 0.00013843 | s
| Standard dev |  | 0.00011057 | s
Difference between trigger signal and peak of NMR signal | Standard dev of trigger time (operational) |  | 7.5736e-05 | s
| Standard dev of peak time |  | 7.4817e-05 | s
| Average difference between trigger and peak time |  | 0.00013587 | s | 0.10285
| Standard dev |  | 0.00010646 | s | 0.080589
PARABOLA FIT | Standard dev of trigger time (spare) |  | 7.4877e-05 | s
| Standard dev of peak time |  | 7.5128e-05 | s
| Average difference between trigger and peak time |  | 0.00016619 | s | 0.12581
| Standard dev |  | 0.00010607 | s | 0.080295